## Slide 1: Machine Learning analysis for absorption measurements correction schemes – A test study
Jesús Yus-Díez1, Jorge Pérez2, Luka Drinovec1,3, Lucas Alados-Arboledas4, Gloria Titos4 and Griša Močnik1,3
1Center for Atmospheric Research, University of Nova Gorica, 5270 Ajdovščina, Slovenia
2Nextail Labs SLU
3Haze instruments d.o.o.
4Andalusian Institute for Earth System Research (IISTA-CEAMA), University of Granada, Granada, Spain

## Slide 2: MEGLICA: ML for Earth Sciences within SMASH
MEGLICA
Machine lEarning applications for constraining the warminG effects of aerosoL partIcles on ClimAte
4.2. Role of anthropogenic and natural aerosols and their contribution to climate warming (UNG-CAR)

## Slide 3: Introduction: Aerosol contribution to climate warning
Still high uncertainty of the effect of aerosol particles on climate
BC and light absorbing aerosols (LAA) are very important due to their warming capacity
(LAA) warm the atmosphere and modify the thermodynamic structure, introducing indirect & semidirect climate changes (cloud formation, precipitation rates, etc.)
socimage.net
Environ. Defense Fund
CO2
BC (direct)
atmosphere
surface
+0,6
+0,9 W/m2
+1,6 W/m2
+1
+2,6
-1,7
BC total:
+1,1 ± 0,7 W/m2
Bond et al., 2013

## Slide 4: Main Objective
Lack of an established reference absorption measurement & large influence of artefacts
Novel correction technique based on a reference absorption instrument and laboratory experiments
&
Application of advanced corrections using both classical and ML

## Slide 5: Methodology: Measurements
AMBIENT MEASUREMENTS
Size-specific correction for different backgrounds via a ML model
Site-specific corrections for Filter Photometers measurements via a reference measurements
PTAAM
AE33
+
Physical aerosol characterization:
Scattering
Particle Size Distribution

## Slide 6: Methodology: Ambient measurement campaign
Summer 2023 campaign over Granada, Spain
Period: June - September 2023
Collocated measurements:
  PTAAM2𝛌: reference absorption
  AE33/MAAP: Attenuation & absorption
  Nephelometer Aurora 3000: Integrating scattering and backscattering
  APS: Particle size distribution
Events:
  Background pollution
  Wildfire particles
  Saharan Dust oubreaks

## Slide 7: Methodology: Compensation algorithms
Filter Photometers, especially AE33, suffer from cross-sensitivity to scattering of the particles, that is mainly due to:
  Capacity to scatter light, described through (SSA)
  Their penetration capacity, described by their size
Problem
Drinovec et al. (2022)
Yus-Diez et al. (2021)
Notes: Data availability: Not all the stations have all the measurements required to perform compensation algorithms and reference measurements
Know-how: It requires skilled and knowledge to perform compensations so that the data keeps its physical meaning

## Slide 8: Methodology: Compensation algorithms
Data availability
Reference absorption measurements
Know-how
Traditional approach limitations
Cover for the lack of data availability by allowing correction using either Particle size and/or scattering.
Provide an easy to use algorithm that data provider can use to correct data
ML algorithm objective
Notes: Data availability: Not all the stations have all the measurements required to perform compensation algorithms and reference measurements
Know-how: It requires skilled and knowledge to perform compensations so that the data keeps its physical meaning

## Slide 9: Methodology: Compensation algorithms
Correct Attenuation
Collocated measurements:
  Reference absorption objective variable
  Attenuation prediction variable(s)
  Scattering constraining variables*
  Particle size dist. constraining variable*
date | xobjective | xcor | x1 | xi | xn
 | … | … | … | … | …
 | … | … | … | … | …
 | … | … | … | … | …
 | … | … | … | … | …
 | … | … | … | … | …
 | … | … | … | … | …
 | … | … | … | … | …
date
constraining variables
* Not always both available

## Slide 10: Methodology: Machine Learning
A multi-linear regression, as baseline model
A Machine Learning model: gradient boosting ensemble of decision trees.
Evaluate the effect of data availability of physically descriptive variables on the correction
We introduce the constraining variables as features, to improve the model training
Notes: Given the multiple cross-sensibility, two main approaches are used

## Slide 11: Methodology: Machine Learning
A more advanced ML model: gradient boosting ensemble of decision trees.
A decision tree:
  Looks for the feature (variable) that best splits the data
  Repeat for n times
  Each leaf is assigned a value which is the average of the target variable considering only the leaf’s points
  Once trained, we just need to follow the new point until the last leaf and use the value as prediction
E.g.:, for a measurement of batn_950 = 1, will end up in the 5th leaf, and will be assigned a prediction of b_abs_450 = 5.306
We combine 100 decision trees, where each one is trained on the errors of its predecessor tree (Boosting).
Notes: A more advanced Machine Learning model: gradient boosting ensemble of decision trees.
A decision tree:
  looks for the feature (variable) that best (according to some criterion, e.g., R^2) splits the data (branches)
  repeats for a number of times
  each leaf is assigned a value which is the average of the target variable considering only the leaf’s points
  Once trained, we just need to follow the new point until the last leaf and use the value as prediction
  For example, a new measurement with batn_950.jy = 1, will end up in the 5th leaf, and will be assigned a prediction of babs_450_v2 = 5.306
We combine 100 decision trees, where each one is trained on the errors of its predecessor tree (Boosting).

## Slide 12: Results: ML vs mlr vs traditional correction
w/o scattering & PSD
w PSD & w/o scat
w scattering & w/o PSD
w PSD & scattering

## Slide 13: Results: ML vs mlr vs traditional correction
w PSD & w/o scat
w PSD & scattering

## Slide 14: Results: Redo analysis for 808 nm
w PSD & w/o scat
w scattering & w/o PSD
w PSD & scattering
Despite the higher error than can be introduced at this wavelength due to the higher noise-to-signal ratio, ML still provide a good correction when all options available, MLR is noisier

## Slide 15: Future steps
  Within AGORA:
    Expand analysis for all all available wavelengths
    Check if the model predicts well using the traditional correction as a reference, for extrapolation purposes of the model to other datasets
  EU data:
    Check if the model predicts well using the traditional correction as a reference for other EU stations
    Retrain model using data from other EU stations, amplifying the background variability input into the model
    Use of particle source and chemical composition to improve prediction, if available, test whether it improves the outcome
  Check other ML algorithms if necessary, possible use of synthetic data generation to fill data gaps

## Slide 16: Conclusions
Both MLR and Gradient Boosting algorithms provide a high-quality correction of the absorption data
Gradient Boosting outperforms MLR, especially under scenarios where scattering is not available
However, a more simple MLR gives a similar result under scenarios with higher data availability
Since the exploratory analysis has provided good results, application to datasets across Europe to be accomplished in the next step

## Slide 17
This project has received funding from the European Union's Horizon Europe research and innovation programme under the Marie Sklodowska-Curie grant agreement No. 101081355.
The operation (SMASH project) is co-funded by the Republic of Slovenia and the European Union from the European Regional Development Fund.
THANK YOU FOR YOUR ATTENTION!

## Slide 18: References
Yus-Díez, Jesús, et al. "Determination of the multiple-scattering correction factor and its cross-sensitivity to scattering and wavelength dependence for different AE33 Aethalometer filter tapes: a multi-instrumental approach." Atmospheric measurement techniques 14.10 (2021): 6335-6355.
Drinovec, Luka, et al. "A dual-wavelength photothermal aerosol absorption monitor: design, calibration and performance." Atmospheric Measurement Techniques 15.12 (2022): 3805-3825.

## Slide 19: Introduction: Aerosol contribution to climate warning
Net cooling effect on Earth’s climate
Overall small reduction of the warming effect of the greenhouse gases
High uncertainty!
-2.0
-1.0
-0.5
0.0
-1.5
ERF (Wm-2)
Adapted from IPCC (2021)
ERF (Wm-2)
0.0
-1.5
1.5
Direct
Cloud interaction